|  | Government College of Engineering Station Road, Osmanpura, Aurangabad – 431 005 “In Pursuit of Global Competitiveness” Phone : (0240) 2366101 Fax : (0240) 2332835 E-Mail – principalgeca@yahoo.com, Web – http://www.geca.ac.in |
| --- | --- |

Corrigendum No. 1

Invitation No.:- TEQIP-II/2016/MH2G07/Shopping/249 Dt. 18-Nov-2016

Name of Item: Vector Network Analysis

Due to some Technical changes in specification:

As per above reference the technical specification has been modified hence requesting you to consider the modified specification as per given below.

Annexure-I

| Sr.No | Item Name | Specifications |
| --- | --- | --- |
| 1 | Vector Network Analysis | Spectrum Analyzer Frequency Range: 9kHz to minimum 25GHz and above, usable to 5 Khz Frequency Reference Accuracy: ± 2 ppm Resolution Bandwidth: Zero span: 10 Hz to 5 MHz Non-zero span: 1Hz to 5MHz Selectivity (-60 dB/ -3 dB): 4:01 Video Bandwidth: 1Hz to 5MHz Stability (Phase noise) at CF = 1 GHz 10 kHz offset: -111 dBc/Hz ( typical) 30 kHz offset: -108 dBc/Hz, (typical) 100 kHz offset: -104 dBc/Hz, (typical) 1 MHz offset: -113 dBc/Hz, (typical) Trace Points: 101, 201, 401, 601, 801, 1001 Amplitude Measurement Range: Displayed average noise level (DANL) to +20 dBm Input Attenuator Range: 0 to 30 dB Maximum safe Input Amplitude: +27 dBm Displayed Avg Noise Level: -133 dBm with pre-amp off @< 7GHz -149 dBm with pre-amp on @< 7GHz Total Amplitude Accuracy: ± 0.6 dB typical Reference Level Range: -170 dBm to +30 dBm Second Harmonic Distortion: < -80 dBc, +50 dBm SHI (nominal) Residual Responses: -110 dBm (nominal) Traces: 4 traces, data/max/average/min Detectors: Normal, positive peak, negative peak, sample, average Markers: Marker types: Normal, noise marker, delta, band power Number of markers or delta markers: 6 Marker functions: Peak, next peak, peak left, peak right, marker to center, minimum search Marker Table Tune: AM/FM General Specifications: Connector Type: 3.5mm (f) , 50 Ω Display: 6.5” transflective, color VGA LED backlit 640 x 480 with anti-glare coating Adaptive display based on ambient light conditions like outdoor sun, outdoor dusk, night vision etc. Speakers: Built in speakers and 3.5mm headphone out Interface: 2 x USB 2.0; 1 x mini USB; 1 x LAN, 1 x mini SD slot Remote Software: Yes , available EMC Field strength immunity: 10 V/m @<26.5 GHz Internal Memory: minimum 4GB, up to 1000 traces Save Formats: State, State+Trace, Picture, .csv Environmental: Operating temperature: -10 ⁰C to +55 ⁰C Humidity: 95% at 40 °C Mil Standard: MIL-STD-810G Weight: Less than 4 kg including battery Power: External AC power adapter Battery Support: Lithium Ion, Rechargeable and Recyclable. Battery operating time: More than 3 hours Cable Antenna Tester Frequency Range: 30KHz to minimum 25GHz and Above Frequency Reference Accuracy: ± 2 ppm Measurement Points: 101, 201, 401, 601, 801, 1001, 1601, 4001, 10001 Directivity: > 40dB Output Power Range High Power: > -2 dBm @ 1 GHz to 9GHz > -6dBm @ 9 to 18GHz > -12dBm @ 18 to 26.5GHz Low Power: < -45 dBm @ > 300 kHz to 26.5 GHz Measurement -Distance-to-fault (dB) -Return loss -VSWR -Distance-to-fault & Return Loss combine -Cable loss (1-port) - insertion loss(2-port) -distance-to-fault (linear or Rho) Immunity to Interference: +16 dBm (nominal) Maximum safe input Level: +27 dBm Vector Network Analysis Full two Port S11,S12,S21,S22 Frequency Range: 30KHz to minimum 25GHz and Above Measurement Points: 101, 201, 401, 601, 801, 1001, 1601, 4001, 10001 IF Bandwidth: 10Hz, 100Hz, 300 Hz, 1 kHz, 3 kHz, 10 kHz, 30 kHz, 100KHz Dynamic Range: > 95 dB @ 300kHz to 9GHz ≥ 91dB @ 9 to 14 GHz ≥ 14 to 18 GHz 90dB , ≥ 18 to 20 GHz 87dB , > 25 to 26.5 GHz 65dB Calibration Types: Mechanical Cal Output Power: High Power: > -2 dBm @ 1 GHz to 9GHz > -6dBm @ 9 to 18GHz > -12dBm @ 18 to 26.5GHz Low Power: < -45 dBm @ > 300 kHz to 26.5 GHz Measurement: VNA S- parameter, Mixed-mode S parameter Port Extension: Yes at both the ports S11 and S21, 30 kHz to 26.5 GHz, enhanced response cal, 100 kHz IF bandwidth, 1001 points: 483 μs /pt |

Principal

Govt. College of Engineering,

Aurangabad